[image: image1.jpg]Fin Dasoription_D-sub-15p female

85-254VAC SO0z ANALOG oUTPUT

Exhaust 02 & C02 Gas Analyzer
Model: LKN-2000-03 TP

15YAC 28 N

230vc 1A




EXHAUST O2 AND CO2 ANALYZER                         LKM2000A


[image: image2.jpg]LED DislnGssn

iaas

CHANNEL SELECT

-~ LKM2000T 34
Kocss)Bosnicr 025 o1 oas et
51 sonain €3 s

Pin Description

s




[image: image3.jpg]LKM2000
LOKAS AUTOMATION
Gas Outet

@ Normal Flow
@ Measuring Flow

Messuing Gas #1

Fermentor

Flon

Ad,
0.3~ 0.7L/MIN

0.
Sensor





[image: image4.jpg](%) uonoesy sjow 05 100 ——

0 - I o - a

210
208
206
204

(%) uonoesy sjow °0 1NO ——

202

144 168 182 216 240 28B4
Culture time (hr)

120




[image: image5.jpg]= O®®®® [LaLoiinn. st o
=4 5 Boeooo

CABEB





[image: image6.png]



[image: image7.png]INTERNATIONAL




[image: image8.jpg]



[image: image9.jpg]PIEES Szt ool
W;Wﬁ
feow [ ood] %
FioR — |siozomovea|





[image: image10.jpg]



[image: image11.jpg]



[image: image12.jpg]EEEFE]

AMEolS e ek EAESH
1 [EE [5¢ [ o00[ 3000 [ o.00[ 3000
2 [1-c02 [% [ ooo[ 3000 o00] 30.00
3 202 [% [ ooo[ 3000 o.00] 3000
4 [2C02 [% [ ooo[ 3000 o00] 30.00
5 [3-02 [% [ ooo[ 3000 o.00] 30.00
6 [3-C02 [% [ ooo[ 3000 o00] 30.00
7 [+02 [% [ ooo[ 3000 o.00] 3000
8 [4C02 [% [ ooo[ 3000 o00] 30.00
9 [5-02 \% [ ooo[ 3000 [ 000 30.00
10[5-C02 [ ooo[ 3000 [ oo0f 30.00

1 [FI_OUR_in ((1.571.3°0.2098)/298)*1000
v2[F1_OUR_out (1.5*1"(ch17100))1303)1 000

v3[FiOUR |

[(73200073)*(chg-ch10))*1000

\ZIEEN |





[image: image13.jpg]L=[ofx]

2-02
2-c02

]
00/00:15:02.18

[Tl
00/00:21:19.43

B4-02
4-c02

s

B HAZL B
W’W]ﬁ
oz [ 207 %
pcoa [ 000 %
o2 | 207 %
Bcoz [ o00] %
o2 [ 207 %
c2 [ o00] %
o2 [ o000] %
[(-co2 [ 000 %
A 2 \ 0.00





[image: image14.png]



[image: image15.png]DEE BIE 210 S50 AN0Q) =STH

Dl(E| S al oemo]

OnlineDataFileV2.0

Monitoring Start = 1888/07/04 21:04:51.370
Logging  Start = 1989/07/04 21:07:05,450
Logaing - Monitoring = 02 DAIZF 22 14.080%
T3d080 Lol E

10 0
IRRRNRERRRRE]
0000 0000 0000 0000 0000 0000
5000 5000 5000 6000 5000 6000
0000 0000 0000 0000 0000 0000
30000 30000 30000 30.000 30000 30000
ZIHAI2F ZI#AIZHE2| E) pu )
00/00:06:23.02 383024 2088
00/00:11:21.72 681729 21.22
00/00:16:20.76 980764 21.21
00/00:21:19.85 1279854 21.21
00/00:26:19.32 1579329 21.21
00/00:31:19.13 1879134 21.22
00/00:36:18.66 2178664 21.20
SBLIY A iz
“ |

S2HAL,

0000
5000
aom
30000
2
0o
Jikin]
Jikin]

0000

5000

000
30,000

3528528580








   





* Do not installed  Fermentor E _ Analog output signal





�




















CH4








�





Rear panel over view


 











Front panel over view


 











CH2








CH1











CH3








CH5








�





�








�





SPECIPICATION





구 분�
O2�
CO2�
�
측정방식�
Electro-chemical�
NDIR�
�
측정범위�
0 ~ 30% vol.�
0 ~ 30% vol.�
�
Accuracy�
+- 0.5% Full scale�
+- 0.5% Full scale�
�
Stability�
+- 1.0% Full scale�
+- 1.0% Full scale�
�
사용온도,습도�
0 ~ 40 °C, 90RH%�
0 ~ 40 °C, 90RH%�
�
Gas flow rate�
300 ~ 700 sccm�
300 ~ 700 sccm�
�
Ext. Outputs�
0 ~ 5vdc, 4 channels�
0 ~ 5vdc, 4 channels�
�
시리얼 포트�
RS232C�
�
Input power�
90 ~ 230vac, 50/60 Hz�
�
Dimension�
LKM2000-03A :  435(W), 360(D), 132.5(H)�
�






�  �





OVER VIEW


LKM2000A는 Fermentation용으로 특별히 설계된 Exhaust Oxygen(O2) and Carbon Dioxide(CO2) Gas monitoring 기기입니다. 설치 및 사용법이 간편하고, Scan방식으로 Multi system (최대5대) operation을 지원하며,  Recorder 출력 및 Windows based monitoring software가 제공됩니다. 


  








APPLICATION


█Fermentation process





FEATURES


█ NDIR sensor determines Carbon Dioxide concentration


█ Electro Chemical sensor determines Oxygen concentration


█ Designed for Use with any brand fermentor


█ High accuracy and long term stability


█ Built - In Gas sampling pump provides regulated delivery of sample gasses


█ Built - In Flow meter enables precise control of the sample gas flow rate


█ Multi Channel(max. 5ch) Operation Unit 


█ Software Calibration by using Air and CO2 reference gas


█ Linearized Analog output of Data to a Recorder or PLC interface ( 0 - 5Vdc.Standard)


█ Windows based Monitoring Software (RS232)


█ The Compact system is Bench-mounting or 19" standard mounting rack 











  





�





RAW Data Display 창


�





Data logging


�





자유로운 


Data Log간격 지정


�





5CH의 O2 & CO2


Reading CH 및 


간단한 연산채널 제공


�








�








Windows based Monitoring Software


█ PC 사양


   CPU: 펜티엄 이상


   OS :  Windows 98, NT, 2000


   Memory: 128MB


█ Program Size : 181KB

















�





Typical Sensitivity to 0 to 30% vol. O2,CO2





ANALOG OUTPUT SIGNAL (Vdc)





O2,CO2 CONCENTRATION (% vol.)





0                2.5                5.0





30

















15 

















0 








Recorder  


SR1000





15pin, femeil D-sub connector


connect to "ANALOG OUPUT" connector at LKM2000-03





CH1 + 


CH1 -


CH2 + 


CH2 -


CH3 + 


CH3 -


CH4 + 


CH4 -


CH5 + 


CH5 -


CH6 + 


CH6 -


CH7 + 


CH7 -


CH8 + 


CH8 -























Pin 7  : Fermentor D _ O2  Analog output signal(CH7)


Pin 8  : Fermentor D _ CO2 Analog output signal(CH8)


Pin 12 : Signal ground








Pin 5  : Fermentor C _ O2  Analog output signal (CH5)


Pin 6  : Fermentor C _ CO2 Analog output signal (CH6)


Pin 11 : Signal ground








Pin 3  : Fermentor B _ O2  Analog output signal(CH3)


Pin 4  : Fermentor B _ CO2 Analog output signal(CH4)


Pin 10 : Signal ground





Pin 1  : Fermentor A _ O2  Analog output signal (CH1)


Pin 2  : Fermentor A _ CO2 Analog output signal(CH2)


Pin 9  : Signal ground





CH1/ Measuring Gas Inlet


CH2


CH3


CH4


CH5








Fermentor A


Fermentor B


Fermentor C 


Fermentor D


Fermentor E








LKM2000-MC


Back side





Recorder Interface


 














Excel Link을 통한 자유로운 Data Editing 기능


�











(주) 로카스           Tel: 042-863-8301,2,3 Fax: 042-863-8304  lokas@lokas.co.kr  www.lokas.co.kr


